## Sheet1
Field | BIO
DIC: Variable abbreviation in data files | From data
DIC: Variable unit | From data or DATA_TYPE_SEQ
DIC: Observation type | profile
DIC: Measured or calculated | measured
DIC: Calculation method and parameters | Pulled from conversion equation in DATA_TYPE_SEQ
DIC: Sampling instrument | Pulled from GEAR_TYPE_SEQ
DIC: Analyzing instrument | SOMMA with UIC 5011 coulometer
DIC: Detailed sampling and analyzing information | Seawater samples were collected from standard depths in 500 mL borosilicate glass reagent bottles. Five mL of water was removed to allow room for thermal expansion and the sample was preserved within 30 minutes of collection with 0.1 mL of a mercuric chloride saturated solution, then sealed with Apiezon M grease and the ground glass stoppers secured in place with rubber bands. The samples were stored at room temperature until on-shore laboratory analysis. The samples were analysed for DIC using a SOMMA sample handling system in conjunction with a coulometer (UIC Inc.) to quantify total CO2 purged from the acidified sample. Measurements of a Certified reference material seawater (CRMs, Scripps Oceanographic Institution) were used for calibrating DIC with each batch of 20 samples being bracketed by duplicate CRM measurements.
DIC: Field replicate information | "Average of replicates" if AVERAGED_DATA == Y, else left blank
DIC: Standardization technique description | Referenced to duplicate measurement of CRM prior to, and following, each daily batch of DIC samples
DIC: Frequency of standardization | To bracket each batch of 20 samples
DIC: CRM manufacturer | Andrew Dickson, Scripps Oceanographic Institute
DIC: Batch number | According to cruise. See OCS Archive>Tracer>Data>Cruise data summary.xls
DIC: Poison used to kill the sample | Mercuric chloride saturated solution
DIC: Poison volume | 100 uL per 500 ml sample bottle = 0.02%
DIC: Poisoning correction description | None
DIC: Uncertainty | 0.15%
DIC: Data quality flag description | Standardize to WOCE
DIC: Method reference (citation) | Dickson, A.G., Sabine, C.L. and Christian, J.R. (Eds.) 2007. Guide to Best Practices for Ocean CO2 Measurements. PICES Special Publication 3, 191 pp., https://cdiac.ess-dive.lbl.gov/ftp/oceans/Handbook_2007/Guide_all_in_one.pdf
DIC: Researcher Name | Kumiko Azetsu-Scott
DIC: Researcher Institution | Bedford Institute of Oceanography
TA: Variable abbreviation in data files | From data
TA: Variable unit | From data or DATA_TYPE_SEQ
TA: Observation type | profile
TA: Measured or calculated | measured
TA: Calculation method and parameters | Pulled from conversion equation in DATA_TYPE_SEQ
TA: Sampling instrument | Pulled from GEAR_TYPE_SEQ
TA: Analyzing instrument | In-house automated sampler with Metrohm Titrando dosimat and Tiamo software
TA: Type of titration | Potentiometric multi point titration
TA: Cell type (open or closed) | Open
TA: Curve fitting method | Gran point
TA: Detailed sampling and analyzing information | Seawater samples were collected from standard depths in 500 mL borosilicate glass reagent bottles. Five mL of water was removed to allow room for thermal expansion and the sample was preserved within 30 minutes of collection with 0.1 mL of a mercuric chloride saturated solution, then sealed with Apiezon M grease and the ground glass stoppers secured in place with rubber bands. The samples were stored at room temperature until on-shore laboratory analysis. TA in the sample was determined using an open cell automated potentiometric titration with a Metrohm Titrando dosing unit controlled by Metrohm Tiamo software. The multi-point titration, using 0.1N HCl titrant containing 35 g of sodium chloride per litre of acid, was performed in a temperature controlled flask held at 25 °C with Gran endpoint determination. Measurements of a Certified reference material seawater (CRMs, Scripps Oceanographic Institution) were used for calibrating TA with each batch of 20 samples being bracketed by duplicate CRM measurements.
TA: Field replicate information | "Average of replicates" if AVERAGED_DATA == Y, else left blank
TA: Standardization technique description | Referenced to duplicate measurements of a CRM before and after each daily batch of samples
TA: Frequency of standardization | To bracket each batch of 20 samples
TA: CRM manufacturer | Andrew Dickson, Scripps Oceanographic Institute
TA: Batch Number | According to cruise. See OCS Archive>Tracer>Data>Cruise data summary.xls
TA: Poison used to kill the sample | Mercuric chloride saturated solution
TA: Poison volume | 100 uL per 500 ml sample bottle = 0.02%
TA: Poisoning correction description | None
TA: Magnitude of blank correction
TA: Uncertainty | 0.25%
TA: Data quality flag description | Standardize to WOCE
TA: Method reference (citation) | Dickson, A.G., Sabine, C.L. and Christian, J.R. (Eds.) 2007. Guide to Best Practices for Ocean CO2 Measurements. PICES Special Publication 3, 191 pp., https://cdiac.ess-dive.lbl.gov/ftp/oceans/Handbook_2007/Guide_all_in_one.pdf
TA: Researcher Name | Kumiko Azetsu-Scott
TA: Researcher Institution | Bedford Institute of Oceanography
pH: Variable abbreviation in data files | From data
pH: pH scale | Total
pH: Observation type | profile
pH: Measured or calculated | measured
pH: Calculation method and parameters | Pulled from conversion equation in DATA_TYPE_SEQ
pH: Sampling instrument | Pulled from GEAR_TYPE_SEQ
pH: Analyzing instrument | Agilent 8454 spectrophotometer with 10 cm pathlength cell with Apollo Instruments autosampler
pH: Temperature of measurement | 25 degrees Cecius
pH: Detailed sampling and analyzing information | Seawater samples were collected from the rosette in screw cap 200 mL borosilicate glass bottles which were then refrigerated for up to 2 hours prior to analysis. Just prior to measurement, the bottles were equilibrated to 25 degrees C in a water bath then transferred into temperature controlled 10 cm quartz cell in an Agilent 8454 spectrophotometer. Following a blank measurement, 25 mL of a 3 milli-molar solution of purified m-cresol purple was added to the sample using the Apollo autosampler to mix the dye solution. Absorbances were measured at 434, 578 and 730 nm and pH determined according to the method described in SOP 6b of Guide to Best Practices for Ocean CO2 measurements (Dickson et al., 2007). The equation of Liu et al. (2011) was used to calculate pH values and a double dye correction was applied.
pH: Field replicate information | "Average of replicates" if AVERAGED_DATA == Y, else left blank
pH: Standardization technique description | Referenced to duplicate measurements of a CRM before and after each daily batch of samples
pH: Frequency of standardization | Each 40 samples
pH: pH values of the standards | According to CRM batch. Calculated from DIC and TA using CO2.SYS
pH: Temperature of standardization | 25 degrees Cecius
pH: Temperature correction method | None
pH: at what temperature was pH reported | 25 degrees Cecius
pH: Uncertainty | 0.001
pH: Data quality flag description | Standardize to WOCE
pH: Method reference (citation) | Dickson, A.G., Sabine, C.L. and Christian, J.R. (Eds.) 2007. Guide to Best Practices for Ocean CO2 Measurements. PICES Special Publication 3, 191 pp., https://cdiac.ess-dive.lbl.gov/ftp/oceans/Handbook_2007/Guide_all_in_one.pdf
pH: Researcher Name | Kumiko Azetsu-Scott
pH: Researcher Institution | Bedford Institute of Oceanography
pCO2A: Variable abbreviation in data files | My understanding is that we should only have pCO2D, PLEASE CORRECT ME IF I'M WRONG
pCO2A: Variable unit
pCO2A: Observation type
pCO2A: Measured or calculated
pCO2A: Calculation method and parameters
pCO2A: Sampling instrument
pCO2A: Location of seawater intake
pCO2A: Depth of seawater intake
pCO2A: Analyzing instrument
pCO2A: Detailed sampling and analyzing information
pCO2A: Equilbrator type
pCO2A: Equilibrator volume (L)
pCO2A: Vented or not
pCO2A: Water flow rate (L/min)
pCO2A: Headspace gas flow rate (L/min)
pCO2A: How was temperature inside the equilibrator measured .
pCO2A: How was pressure inside the equilibrator measured.
pCO2A: Drying method for CO2 gas
pCO2A: Manufacturer of the gas detector
pCO2A: Model of the gas detector
pCO2A: Resolution of the gas detector
pCO2A: Uncertainty of the gas detector
pCO2A: Standardization technique description
pCO2A: Frequency of standardization
pCO2A: Manufacturer of standard gas
pCO2A: Concentrations of standard gas
pCO2A: Uncertainties of standard gas
pCO2A: Water vapor correction method
pCO2A: Temperature correction method
pCO2A: at what temperature was pCO2 reported
pCO2A: Uncertainty
pCO2A: Data quality flag description
pCO2A: Method reference (citation)
pCO2A: Researcher Name
pCO2A: Researcher Institution
pCO2D: Variable abbreviation in data files | From data
pCO2D: Variable unit | From data or DATA_TYPE_SEQ
pCO2D: Observation type | profile
pCO2D: Measured or calculated | measured
pCO2D: Calculation method and parameters | Pulled from conversion equation in DATA_TYPE_SEQ
pCO2D: Sampling instrument | Pulled from GEAR_TYPE_SEQ
pCO2D: Analyzing instrument | SRI 8610C gas chromatograph with flame ionization detector and methanizer
pCO2D: Storage method | Vials stored in refrigerator
pCO2D: Seawater volume (mL) | 150 mL
pCO2D: Headspace volume (mL) | 11 mL
pCO2D: Temperature of measurement | 22 degrees Celcius
pCO2D: Detailed sampling and analyzing information | Seawater samples were collected in 160 mL volume crimp seal serum bottles, allowing the bottle to overflow by 3 volumes. The sample was immediately stabilised by the addition of 50 mL of saturated mercuric chloride solution, crimp sealed with a butyl rubber septum then stored in a refrigerator. Before analysis, a 11 mL headspace of 400 ppm CO2 in zero air was introduced into the bottle which was then thermally equilibrated at 22 °C for one hour. The bottle was then shaken vigorously for 8 minutes and the headspace displaced by a brine solution, flushing the sample loop of an SRI gas chromatograph equipped with a methaniser and flame ionisation detector. CO2 peaks were calibrated by injections of primary standard gas mixtures (Air Liquide) having CO2 mixing ratios of 397.3, 797.2, 1201 and 2017 ppm.
pCO2D: Field replicate information | "Average of replicates" if AVERAGED_DATA == Y, else left blank
pCO2D: Manufacturer of the gas detector | SRI instruments
pCO2D: Model of the gas detector | 8610 GC-FID with methanizer
pCO2D: Resolution of the gas detector | 0.1 uatm
pCO2D: Uncertainty of the gas detector | 1 uatm
pCO2D: Standardization technique description | Calibration plot made up from injections of primary gas standards, 400 ppm, 800 ppm, 1200 ppm
pCO2D: Frequency of standardization | Twice daily, before and after each batch of samples
pCO2D: Temperature of standardization | Sample loop held at 40 degrees
pCO2D: Manufacturer of standard gas | Air Liquide
pCO2D: Concentrations of standard gas | 397.3 ppm, 797.2 ppm, 1201 ppm
pCO2D: Uncertainties of standard gas | 1%
pCO2D: Water vapor correction method | Dickson et al., 2007, Guide to best practices for ocean CO2 measurements, Sop 4, section 8.3
pCO2D: Temperature correction method | None
pCO2D: at what temperature was pCO2 reported | 22 degrees Celcius
pCO2D: Uncertainty | 0.25%
pCO2D: Data quality flag description | Standardize to WOCE
pCO2D: Method reference (citation) | Dickson, A.G., Sabine, C.L. and Christian, J.R. (Eds.) 2007. Guide to Best Practices for Ocean CO2 Measurements. PICES Special Publication 3, 191 pp., https://cdiac.ess-dive.lbl.gov/ftp/oceans/Handbook_2007/Guide_all_in_one.pdf
pCO2D: Researcher Name | Kumiko Azetsu-Scott
pCO2D: Researcher Institution | Bedford Institute of Oceanography
Author | Carrie-Ellen Gabriel; Darlene Childs